## Sheet1
Job Record Number | Customer Name | Distributer/Seller | Model Number | Model Variant | Serial Number | Original Software Version | Original GUI Version | Manufacture Date | Shipping Date | Individual | Notes/Comments
- | Inventory | - | S79.01 | 2019-06-01 | - | - | - | - | -
- | Malaysia Training Tool | Enersol Pty Ltd. | S78.02 | 2019-04-01 | - | S78.02.2.0.6 | S78.02.2.0.1 | Re 8/09/2020 | 2020-09-22 |  | Remanufactured from software exchange unit. Model Variant Apr-19 with updated software. To be used as a training aid for Enersol Malaysia staff. Has the serial number S78.02.0018 but Hung says to treat like it has none in case of serial number clash with another customer.
- | Enersol Sdn. Bhd. | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0008 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-01-31 | -
- | Enersol | Enersol Pty Ltd. | ELTP.01 |  | ELTP.01.0001 | [5 empty] | Prototype ELT+ in the Wetlab
- | Enersol Sdn. Bhd. | Enersol Pty Ltd. | S77.01 | 2017-03-01 | S77.01.0002 | S77.01.1.0.1 | S77.01.1.0.0 | 2017-07-23 | 2017-07-25
- | Enersol Pty. Ltd. | Enersol Pty Ltd. | S77.01 | 2018-04-01 | S77.01.0005 | S77.01.1.1.2 | S77.01.1.1.1 | 2018-04-23 | 2018-04-23
- | Inventory | - | S77A.03 | 2018-09-01 | S77A.03.0001 | S77A.3.0.0 | S77A.3.0.0 | 2020-06-11 | -
- | Enersol Sdn. Bhd. | Enersol Pty Ltd. | S78.01 | 2017-03-01 | S78.01.0003 | S78.01.1.0.2 | S78.01.1.1.0 | 2017-07-23 | 2017-07-25
- | Inventory | - | S78.01 | HA Apr-18 | S78.01.0007 | S78.01.1.1.2 | S78.01.1.1.1 | 2018-04-23 | -
- | Inventory | - | S78B.03 | 2020-03-01 | S78B.03.0000 | S78B.3.0.0 | S78B.3.0.0 | - | -
- | Enersol Sdn. Bhd. | Enersol Pty Ltd. | S79.01 | 2017-04-01 | S79.01.0002 | S79.01.1.0.0 | S79.01.1.0.1 | 2017-07-23 | 2017-07-25
- | Enersol Sdn. Bhd. | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0002 | VLT.2.0.0 | VLT.GUI.2.0.0 | 2021-08-16 | Late 2021 |  | Second version of the new VLT device. Shipped to Malaysia with parts to be assembled and used for testing in the Enersol Malaysia Lab.
(Replacment) | Invima |  | S77.02 | 2018-09-01 | S77.02.0017 | S77.2.0.5 | S77.GUI.2.0.4 | 2022-11-08 |  |  | Replacment for Invima S77 (S77.01.0007) which was experiencing issues with the screen. The original device will be returned to Enersol after the next AMC
?? | Enersol Pty. Ltd. (Malaysia) | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0019 | AWPSIT.MC.01.1.1.8 | AWPSIT.GUI.01.1.1.8 | 2022-03-23 |  |  | Internal components (PCB, screen and arduino) for the upgrade of the original AWPSIT.MC.01.0002 which was having consistent issues with freezing.
?? | Enersol Pty. Ltd. (Malaysia) | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0020 | AWPSIT.MC.01.1.1.8 | AWPSIT.GUI.01.1.1.8 | 2023-05-11 |  |  | Replacment internal compoinents for the malaysian device above. Which have had issues during installation.
?? | SPARE |  | ELTP.01 |  | ELTP.01.0003 | ELTP.1.0.1 | ELTP.GUI.1.0.0 |  |  |  | Dual Bank
?? | Enersol Pty. Ltd. (Malaysia) |  | ELTP.01 |  | ELTP.01.0004 | ELTP.1.0.0 | ELTP.GUI.1.0.0 |  |  |  | Single Bank
JR2016113 | Ulinzi Latex | Enersol Pty Ltd. | AWPSIT.MC.01 | 2016-12-01 | AWPSIT.MC.01.0001 | AWPSIT.MC.01.1.0.0 | AWPSIT.MC.01.1.0.0 | 2016-12-20 | 2017-02-23
JR2017009 | Shandong Quality Inspection Centre For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S77.01 | 2017-03-01 | S77.01.0001 | S77.01.1.0.1 | S77.01.1.0.0 | 2017-03-19 | 2017-05-02
JR2017009 | Shandong Quality Inspection Centre For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S78.01 | 2017-03-01 | S78.01.0001 | S78.01.1.0.0 | S78.01.1.0.0 | 2017-03-19 | 2017-05-02
JR2017009 | Shandong Quality Inspection Centre For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2017-02-01 | S79.01.0001 | S79.01.1.0.0 | S79.01.1.0.1 | 2016-02-02 | 2017-05-02
JR2017083 | B.Braun Medical Industries Sdn. Bhd. | Basan the Cleanroom Company | S78.01 | 2017-03-01 | S78.01.0004 | S78.01.1.0.2 | S78.01.1.1.0 | 2017-08-25 | 2017-09-11
JR2017083 | B.Braun Medical Industries Sdn. Bhd. | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0008 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-04-02 | 2019-04-14 |  | 01/04/2019 - Test model for the prefabricated control boxes cut by Plastix Centre. - New protective plastic over Wenchai Labels. - Repurposed DL Medica's intended transducer which had soldered wires on them.
JR2017087 | Whitehouse Analytical Lab | Gage Crib Worldwide, Inc. | S77.01 | 2017-06-01 | S77.01.0004 | S77.01.1.1.1 | S77.01.1.1.1 | 2017-11-03 | 2017-11-09
JR2017087 | Whitehouse Analytical Lab | Gage Crib Worldwide, Inc. | S77.02 | 2018-09-01 | S77.02.0005 |  |  | 2019-01-24 | 2019-02-22
JR2017087 | Whitehouse Analytical Lab | Gage Crib Worldwide, Inc. | S78.01 | 2017-06-01 | S78.01.0006 | S78.01.1.1.1 | S78.01.1.1.1 | 2017-11-03 | 2017-11-09 |  | 9/11/2017 Machine shipped: - 24Vdc label was placed above the power supply socket instead of 12Vdc. New sticker needs to be sent.
JR2017087 | Whitehouse Analytical Lab | Gage Crib Worldwide, Inc. | S78.02 | 2018-09-01 | S78.02.0007 | S78.02.2.0.4 | S78.02.2.0.0 | 2019-01-30 | 2019-02-22
JR2017087 | Whitehouse Analytical Lab | Gage Crib Worldwide, Inc. | S79.01 | 2017-04-01 | S79.01.0003 | S79.01.1.0.0 | S79.01.1.0.1 | 2017-11-07 | 2017-11-09
JR2017090 | Tanzania Food and Drugs Authority | Enersol Pty Ltd. | AWPSIT.MC.01 | 2017-10-01 | AWPSIT.MC.01.0003 | AWPSIT.MC.01.1.0.3 | AWPSIT.MC.01.1.0.0 | 2017-11-22 | 2017-11-28
JR2017091 | Enersol Sdn. Bhd. | Enersol Pty Ltd. | AWPSIT.MC.01 | 2016-12-01 | AWPSIT.MC.01.0002 | AWPSIT.MC.01.1.0.1 | AWPSIT.MC.01.1.0.0 | 2017-07-23 | 2017-09-19 |  | 23/07/2017 - The tank lid has two additional through holes. These holes are being plugged by grub screws as they go straight through the tank. Could cause leaks over time if the seal made by the grub screws erode. - Power distribution board needs the pulldown resistors removed from across the lid sensor and emergency stop button sections. Software updates are needed to accomodate the change to the power distribution board.
JR2017099 | RTI Surgical | Gage Crib Worldwide, Inc. | S78.01 | 2018-04-01 | S78.01.0005 | S78.01.1.1.1 | S78.01.1.1.1 | 2017-11-09 | 2017-11-14
JR2017099 | RTI Surgical | Gage Crib Worldwide, Inc. | S78.02 | 2018-09-01 | S78.02.0005 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-01-24 | 2019-11-04 |  | 14/02/2019 - Updated software to S78.2.0.5 while the machine is still in inventory.
JR2017113 | Saudi Food and Drug Authority | Masareb Gulf Trading | AWPSIT.MC.01 | 2017-12-01 | AWPSIT.MC.01.0004 | AWPSIT.MC.01.1.0.4 | AWPSIT.MC.01.1.0.4 | 2017-12-07 | 2017-12-23 |  | 18/12/2017 - Operation manual December 2017 v2 sent with box. Was a final draft and may be outdated soon. - Packed with styrofoam cutouts. - Parts wrapped in bubblewrap.
JR2018013 | Laboratorio Físico-Mecánico de Dispositivos Médicos y Otras Tecnologías - Instituto Nacional de Vigilancia de Medicamentos y Alimentos - Invima | Enersol Pty Ltd. | S77.01 | 2018-09-01 | S77.01.0007 | S77.01.1.1.3 | S77.01.1.1.1 | 2018-07-26 | 2018-10-18 |  | 4/10/2018 - Updated to S77.02 model
JR2018013 | Laboratorio Físico-Mecánico de Dispositivos Médicos y Otras Tecnologías - Instituto Nacional de Vigilancia de Medicamentos y Alimentos - Invima | Enersol Pty Ltd. | S78.01 | HA Apr-18 | S78.01.0010 | S78.01.1.1.2 | S78.01.1.1.1 | 2018-07-26 | 2018-10-18
JR2018038 | Becton, Dickinson & Company | Gage Crib Worldwide, Inc. | S77.01 | 2018-04-01 | S77.01.0003 | S77.01.1.1.1 | S77.01.1.1.1 | 2018-04-23 | 2018-06-22
JR2018038 | Beckton, Dickinson & Company | Gage Crib Worldwide, Inc. | S77.02 | 2018-12-01 | S77.02.0004 |  |  | 2018-12-19 | 2019-02-08
JR2018038 | Beckton, Dickinson & Company | Gage Crib Worldwide, Inc. | S78.01 | 2017-06-01 | S78.01.0002 | S78.01.1.1.1 | S78.01.1.1.1 | 2018-03-26 | 2018-06-22
JR2018038 | Beckton, Dickinson & Company | Gage Crib Worldwide, Inc. | S78.02 | 2018-09-01 | S78.02.0004 | S78.02.2.0.4 | S78.02.2.0.0 | 2018-12-19 | 2019-02-08
JR2018038 | Beckton, Dickinson & Company | Gage Crib Worldwide, Inc. | S79.01 | 2017-11-01 | S79.01.0004 | S79.01.1.0.0 | S79.01.1.0.1 | 2018-05-23 | 2018-06-22
JR2018039 | Makmal Ujian Fizikal, Lembaga Getah Malaysia, Malaysia | SND Sejagat Enterprise | AWPSIT.MC.01 | 2017-12-01 | AWPSIT.MC.01.0005 | AWPSIT.MC.01.1.0.6 | AWPSIT.MC.01.1.0.4 | 2018-06-07 | 2018-07-05
JR2018045 | SSL Thailand | Enersol Pty Ltd. | AWPSIT.MC.01 | 2017-12-01 | AWPSIT.MC.01.0006 | AWPSIT.MC.01.1.0.6 | AWPSIT.MC.01.1.0.4 | 2018-08-13 | 2018-08-15
JR2018048 | Terumo - Clincal Supply | Enersol Pty Ltd. | S78.01 | 2018-04-01 | S78.01.0008 | S78.01.1.1.2 | S78.01.1.1.1 | 2018-07-09 | 2018-07-27
JR2018055 | Hubei Medical Apparatus and Instruments Quality Supervision and Inspection Center | Jinan Huanzheng Technology & Development Co., Ltd. | S77.01 | 2018-04-01 | S77.01.0006 | S77.01.1.1.2 | S77.01.1.1.1 | 2018-06-14 | 2018-09-04 |  | 14/06/2018 - System was updated last minute to include the pressure vessel as replacement to the flow restrictor (predecessor to the S77.02 Sep-18 model).
JR2018055 | Hubei Medical Apparatus and Instruments Quality Supervision and Inspection Center | Jinan Huanzheng Technology & Development Co., Ltd. | S78.01 | 2018-07-01 | S78.01.0009 | S78.01.1.1.2 | S78.01.1.1.1 | 2018-08-27 | 2018-09-04 |  | 11/07/2018 - This unit contains SBCT-WPSIT Power Hub v1 with the linear voltage regulator instead of the switchmode voltage regulator. - Last minute improvements were made to the system. Contains pressure vessel inside the system (predecessor to the S78.02 Sep-18 model).
JR2018055 | Hubei Medical Apparatus and Instruments Quality Supervision and Inspection Center | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2017-11-01 | S79.01.0005 | RotaryEncoderWithInterrupts_v3.1 | S79.01.1.0.1 | 2018-06-22 | 2018-09-04 |  | 21/06/2018: - Initial software flashed onto the Arduino had communication issues. Rolled back to the older software and the system worked fine.
JR2018075 | Saldanha Rodrigues Ltda | Enersol Pty Ltd. | S77.02 | 2018-09-01 | S77.02.0001 | S77.02.2.0.0 | S77.01.1.1.1 | 2018-10-29 | 2018-11-12
JR2018075 | Laboratório Cepromed | Enersol Pty Ltd. | S77.02 | 2018-09-01 | S77.02.0002 | S77.02.2.0.0 | S77.01.1.1.1 | 2018-10-29 | 2018-11-12
JR2018075 | Saldanha Rodrigues Ltda | Enersol Pty Ltd. | S78.02 | 2018-09-01 | S78.02.0001 | S78.02.2.0.0 | S78.01.1.1.1 | 2018-10-29 | 2018-11-12
JR2018075 | Laboratório Cepromed | Enersol Pty Ltd. | S78.02 | 2018-09-01 | S78.02.0002 | S78.02.2.0.0 | S78.01.1.1.1 | 2018-10-03 | 2018-11-12 |  | 07/09/2018 - System contains an AIR-VAC UV143H vacuum ejector as this device is being sent to a laboratory that falls within the failure range for the Festo VN-10-H-T3-PQ2-VQ2-RO1 vacuum ejector
JR2018080 | Terumo Interventional Systems | Enersol Pty Ltd. | S78.01 | 2018-04-01 | S78.01.0011 | S78.01.1.0.0 | S78.01.1.0.0 | 2017-03-19 | 2018-08-13 |  | 26/05/2017 - The normalizing ratio for calculating the leakage rate is wrong. A software update is needed, new control box is pending. 23/04/2018 - Updated the software and grips. Replaced the entire system. Transducer will be replaced with another calibrated one. 08/08/2018 - Serial number has been replaced from S78.01.0002 to S78.01.0011. Beckton & Dickinson was mistakenly given S78.01.0002. - Horizontal sample support post has been updated with a newer model. - Updated firmware: - GUI vS78.01.1.0.0 -> vS78.01.1.1.1 - Source Code vS78.01.1.0.0 -> vS78.01.1.1.2
JR2018082 | Heilongjiang Institute For Food and Drug Control | Jinan Huanzheng Technology & Development Co., Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0007 | AWPSIT.MC.01.1.0.6 | AWPSIT.MC.01.1.0.4 | 2018-10-18 | 2018-12-14
JR2018106 | Descarpack | Enersol Pty Ltd. | S77.02 | 2018-11-01 | S77.02.0003 | S77.02.2.0.2 | S77.02.2.0.2 | 2018-11-15 | 2019-01-14 |  | 15/11/2018: - S77 Connector junction block faulty production batch. Used loctite on the top exhaust valve mounting screw to plug the breaching hole in the transducer port. Used plugs to fill in volume in the connector junction block as the holes were drilled too deep.
JR2018106 | Descarpack | Enersol Pty Ltd. | S78.02 | 2018-11-01 | S78.02.0003 | S78.02.2.0.3 | S78.02.2.0.0 | 2018-10-15 | 2019-01-14 |  | 15/11/2018: - Reused an old S78 brass connector junction block. - Fitted with a high power venturi tube
JR2018106 | Descarpack | Enersol Pty Ltd. | S79.01 | 2018-11-01 | S79.01.0006 | RotaryEncoderWithInterrupts_v3.1 | S79.01.1.0.1 | 2018-11-14 | 2019-01-14 |  | 14/11/2018: - New Power Hub installed. - Slot was cut to easily align the base clamp mounting screws. - Matting was washed with harsh acetone. Some areas of the blue matting have lost the blue finish and retains some texter markings.
JR2018150 | West Pharmaceutical Services | Gage Crib Worldwide, Inc. | S77.02 | 2018-09-01 | S77.02.0006 |  |  | 2019-01-29 | 2019-02-22
JR2018150 | West Pharmaceutical Services | Gage Crib Worldwide, Inc. | S78.02 | 2018-09-01 | S78.02.0006 | S78.02.2.0.4 | S78.02.2.0.0 | 2019-01-24 | 2019-02-22
JR2019062 | SSL Thailand | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0010 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-05-20 | 2019-07-02
JR2019089 | SSL Thailand | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0009 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-05-20 | 2019-07-02 |  | Fault in machine was discovered when machine was being installed onsite for customer. Serial number or software version could not be read back, machine would lock up when attempting to do so. Arduino, PCB, screen were sent back. On investigation the issue lay in the SD Card. Issue is suspected to be in the brand of SD card, Lexar. Tested working replacement components were sent for installation. Returned PCB was given to RD111a
JR2019090 | Henan Province Medical Instrument Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0011 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-07-26 | 2019-09-12
JR2019090 | Henan Province Medical Instrument Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | S77.02 | 2019-02-01 | S77.02.0007 | S77.02.2.0.3 | S77.02.2.0.2 | 2019-02-27 | 2019-09-12 |  | 27/02/2019: - Drawing specified a 26mm gap between the holes drilled for the connector junction block on the control box. This clearance does not work when using Simon's new white hole plugger; using it would centre the quick connector raising the M8 bolt hole too high to line up with the hole drilled on the control box. Fix was to file material away from the upper side of the M8 hole drilled in the control box, recommend updating the drawing to reduce the gap to 25mm.
JR2019090 | Henan Province Medical Instrument Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | S78.02 | 2019-04-01 | S78.02.0009 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-05-28 | 2019-09-12
JR2019090 | Henan Province Medical Instrument Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2019-06-01 | S79.01.0007 | S79.01.1.0.1 | S79.01.1.0.1 | 2019-06-03 | 2019-09-12
JR2019120 | Togo Medikit Co. Ltd., Hyoga 2nd Factory | R Medical Devices Llc. | S78.02 | 2019-04-01 | S78.02.0010 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-09-27 | 2019-10-31
JR2019122 | Shanghai Testing & Inspection Institute For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S77.02 | 2019-09-01 | S77.02.0008 | S77.02.2.0.4 | S77.02.2.0.4 | 2019-09-12 | 2019-11-22 |  | 15/04/2021 - How does this have S77.2.0.4 GUI if latest version is S77.2.0.3
JR2019122 | Shanghai Testing & Inspection Institute For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S78.02 | 2019-04-01 | S78.02.0011 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-11-11 | 2019-11-22
JR2019122 | Shanghai Testing & Inspection Institute For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2019-06-01 | S79.01.0008 | S79.01.1.0.1 | S79.01.1.0.1 | 2019-10-21 | 2019-11-22
JR2019143 | Shandong Quality Inspection Centre For Medical Devices | Jinan Huanzheng Technology & Development Co., Ltd. | S78A.03 | 2020-03-01 | S78A.03.0001 | S78A.3.0.1 | S78A.03.1.0.1 | 2019-12-05 | 2020-06-16
JR2019148 | DL Medica SA de CV | Gage Crib Worldwide, Inc. | S78H.02 | HA Nov-19 | S78.02.0012 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-11-28 | 2019-11-29
JR2019168 | Bomra Botswana | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0014 | AWPSIT.MC.01.1.1.0 | AWPSIT.MC.01.1.0.8 | 2021-03-30 | 2021-04-12 |  | 31/03/2021 - Decayed rubber bands were left on the u-brackets and sliding bar. Removed and cleaned but there may still be some rubber residue/fragments that couldn't be removed.
JR2019173 | Terumo (Philippines) Corporation | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0014 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-11-25 | 2019-01-31
JR2019178 | Cadence Rhode Island | Gage Crib Worldwide, Inc. | S77.02 | 2018-09-01 | S77.02.0009 | S77.02.2.0.3 | S77.02.2.0.2 | 2019-11-21 | 2019-12-06
JR2019178 | Cadence Rhode Island | Gage Crib Worldwide, Inc. | S78.02 | 2019-04-01 | S78.02.0013 | S78.02.2.0.5 | S78.02.2.0.0 | 2019-11-21 | 2019-12-06
JR2019180 | SSL Shanghai Taicang | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0012 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-11-18 | 2019-12-20
JR2019180 | SSL Shanghai Taicang | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0013 | AWPSIT.MC.01.1.0.7 | AWPSIT.MC.01.1.0.4 | 2019-11-18 | 2019-12-20
JR2020054 | Medline Industries Inc. | Gage Crib Worldwide, Inc. | S78.02 | 2019-04-01 | S78.02.0015 | S78.02.2.0.5 | S78.02.2.0.0 | 2020-02-03 | 2020-05-26
JR2020059 | Ambu Sdn. Bdh. | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0016 | S78.02.2.0.5 | S78.02.2.0.0 | 2020-05-06 | 2020-06-19
JR2020076 | Nephron Pharmaceuticals Corp. | Gage Crib Worldwide, Inc. | S78.02 | 2019-04-01 | S78.02.0017 | S78.02.2.0.5 | S78.02.2.0.0 | 2020-05-12 | 2020-07-08
JR2020104 | B.Braun Medical Industries Sdn. Bhd. | Basan the Cleanroom Company | S77.02 | 2018-09-01 | S77.02.0011 | S77.02.2.0.4 | S77.02.2.0.3 | 2020-08-19 | 2020-10-09 |  | 20/08/2020 - B Braun asked for the ability to change the test duration. V2.0.4 is the introduction of adjustable test duration.
JR2020111 | Lifelong Meditech Ltd. | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0020 | S78.02.2.0.6 | S78.02.2.0.1 | 2021-04-13 | - |  | 07/04/2021 - Machine is a refurb of S78.01.0006
JR2020161 | Terumo (Philippines) Corporation | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0018 | S78.02.2.0.6 | S78.02.2.0.1 | 2020-12-08 | 2021-03-04 |  | 08/12/2020 - RES and GND pins on one of the serial header rows on the back of the touchscreen are bent at odd angles.
JR2021027 | Rwanda Food And Drug Authority | Palin Corporation | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0016 | AWPSIT.MC.01.1.0.8 | AWPSIT.MC.01.1.0.8 | 2021-04-22 | -
JR2021033 | Shenzen Institute For Drug Control, China | Jinan Huanzheng Technology & Development Co., Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0015 | AWPSIT.MC.01.1.1.0 | AWPSIT.MC.01.1.0.8 | 2021-04-06 | 2021-05-13
JR2021033 | Shenzen Institute For Drug Control, China | Jinan Huanzheng Technology & Development Co., Ltd. | S77.02 | 2018-09-01 | S77.02.0010 | S77.02.2.0.4 | S77.02.2.0.3 | 2021-04-13 | 2021-05-13
JR2021033 | Shenzen Institute For Drug Control, China | Jinan Huanzheng Technology & Development Co., Ltd. | S78.02 | 2019-04-01 | S78.02.0019 | S78.02.2.0.6 | S78.02.2.0.1 | 2021-04-13 | 2021-05-13 |  | 16/12/2020 - Refurbed S78.01.0004.
JR2021105 | Central Drug Laboratory, Karachi | USP | VLT.02 | 2021-08-01 | VLT.02.0001 | VLT.2.0.0 | VLT.GUI.2.0.0 | 2021-08-16 | Late 2021 |  | First of the new touch screen VLT devices. Control panels and all other parts were shipped to malaysia to be assembled. A secondary Control box was also shipped for Malaysia to use.
JR2021201 | Gansu Medical Device Inspection and Testing Institute | - | S77.02 | 2018-09-01 | S77.02.0012 | S77.02.2.0.3 | S77.02.2.0.2 | - | -
JR2021201 | Gansu Medical Device Inspection and Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | S78H.02 | HA Apr-19 | S78.02.0022 | S78.02.2.0.6 | S78.02.2.0.1 | - | - |  | 22/11/2021 - Using the S78H parts (S78.01.0007) that was intended for DL medica order that never went through. Parts have been pilfered to other machines during inventory shortage.
JR2021201 | Gansu Medical Device Inspection and Testing Institute | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2019-06-01 | S79.01.0009 | S79.01.1.0.1 | S79.01.1.0.1 | - | -
JR2021203 | Terumo Interventional Systems | R Medical Devices Llc. | S78.02 | 2019-04-01 | S78.02.0021 | S78.02.2.0.6 | S78.02.2.0.1 | 2021-11-08 | - |  | 08/11/2021 - Replacement unit for Terumo's S78.01.0011 22/11/2021 - Using S78.01.0003's front cover and sample testing block. The rest is brand new.
JR2021215 | B.Braun Medical Industries Sdn. Bhd. | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0023 | S78.02.2.0.6 | S78.02.2.0.1 | - | - |  | 08/09/2023 - Ghazaly has replaced the PCB and Festo Valve to resolve a valve issue. I believe the Valve 1 Transistor became damaged during use and was always on. Spare Arduino Supplied with same serial number to be returned to sydney. Keep an eye out for this device to ensure there isnt leaking as ghazaly changed the festo valve.
JR2021215 | B.Braun Medical Industries Sdn. Bhd. | Enersol Pty Ltd. | S78.02 | 2019-04-01 | S78.02.0024 | S78.02.2.0.6 | S78.02.2.0.1 | - | - |  | 22/11/2021 - Using the front panel from S78.02.0021. Only a new back panel was used for the S78.02.0021 job.
JR2021228 | Qindao SGS | - | S77.02 | 2018-09-01 | S77.02.0013 | S77.2.0.4 | S77.GUI.2.0.3 | 2020-05-05 | -
JR2021228 | Jiangsu Center | - | S77.02 | 2018-09-01 | S77.02.0014 | S77.2.0.4 | S77.GUI.2.0.3 | 2020-05-05 | -
JR2021228 | Jiangsu Center | - | S77.02 | 2018-09-01 | S77.02.0015 | S77.2.0.4 | S77.GUI.2.0.3 | 2022-01-12 | -
JR2021228 | Qindao SGS | Jinan Huanzheng Technology & Development Co., Ltd. | S78B.03 | 2020-03-01 | S78B.03.0001 | S78B.3.0.0 | S78B.3.0.0 | - | -
JR2021228 | Jiangsu Center | Jinan Huanzheng Technology & Development Co., Ltd. | S78B.03 | 2020-03-01 | S78B.03.0002 | S78B.3.0.0 | S78B.3.0.0 | - | -
JR2021228 | Jiangsu Center | Jinan Huanzheng Technology & Development Co., Ltd. | S78B.03 | 2020-03-01 | S78B.03.0003 | S78B.3.0.0 | S78B.3.0.0 | - | -
JR2021228 | Qindao SGS | Jinan Huanzheng Technology & Development Co., Ltd. | S79.01 | 2019-06-01 | S79.01.0010 | S79.01.1.0.1 | S79.01.1.0.1 | - | -
JR2022057 | Jiangxi Center of Medical Device Testing | Jinan Huanzheng Technology & Development Co., Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0017 | AWPSIT.MC.01.1.1.6 | AWPSIT.GUI.01.1.1.6 | 2022-03-23
JR2022077 | Medicines Control Authority of Zimbabwe |  | VLT.02 | 2021-08-01 | VLT.02.0003 | VLT.2.0.3 | VLT.GUI.2.0.1 |  |  |  | 1/2 VLT for Zimbabwe (Glove Unit)
JR2022077 | Medicines Control Authority of Zimbabwe |  | VLT.02 | 2021-08-01 | VLT.02.0004 | VLT.2.0.3 | VLT.GUI.2.0.1 |  |  |  | 2/2 VLT for Zimbabwe (Glove Unit)
JR2022128 | Balton Sp z.o.o | Enersol Pty Ltd. | S77.02 | 2018-09-01 | S77.02.0016 | S77.2.0.4 | S77.GUI.2.0.3 | 2022-08-01
JR2022128 | Balton Sp. Z.o.o | Enersol Pty Ltd. | S78B.03 | 2020-03-01 | S78B.03.0004 | S78B.3.1.0 |  | 2022-08-01
JR2022145 | TGA Australia | Enersol Pty Ltd. | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0018 | AWPSIT.MC.01.1.1.7 | AWPSIT.GUI.01.1.1.7 | 2022-03-23
JR2022145 | TGA Australia | Enersol Pty Ltd. | ELTP.01 |  | ELTP.01.0002 | ELTP.1.0.0 | ELTP.GUI.1.0.0 | 2022-10-01 |  |  | First release ELT+ device
JR2022145 | TGA Australia | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0005 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2022-10-01 |  |  | 1 VLT for TGA (Glove Only)
JR2022196 | VISION SCIENTIFIC & ENGG. KENYA LTD | Avantgarde Inc | VLT.02 | 2021-08-01 | VLT.02.0006 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2022-11-01 |  |  | Glove Device
JR2022196 | Female Health Comp (M) S/B | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0008 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2023-02-24 |  |  | Device for assembly in malaysia
JR2022200 | Ministry of Health Ethiopia | Clinton Health Access Initiative | VLT.02 | 2021-08-01 | VLT.02.0007 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2022-11-01 |  |  | Glove Device
JR2022202 | Female Health Co (UK) PLC |  | AWPSIT.FC.01 | 2018-10-01 | AWPSIT.FC.01.0001 | AWPSIT.MC.01.1.1.7 | AWPSIT.GUI.01.1.1.7 | 2022-11-17 |  |  | Female Condom Device
JR2023? | Enersol Malaysia |  | VLT.02 | 2021-08-01 | VLT.02.0009 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2023-02-24 |  |  | Parts for malaysia to assembly a unit for themselves
JR2023006 | Nutricia Pharmaceuticals | Jinan Huanzheng Technology & Development Co., Ltd. | S78B.03 | 2020-03-01 | S78B.03.0005 | S78B.3.1.0 | S78B.GUI.3.0.1 | 2023-02-03
JR2023045 | TSK Laboratory International Japan | R Medical Devices Llc. | S78B.03 | 2020-03-01 | S78B.03.0006 | S78B.3.1.1 | S78B.GUI.3.0.2 | 2023-04-19
JR2023052 | TSK Laboratory International Japan | R Medical Devices Llc. | S77.02 | 2018-09-01 | S77.02.0018 | S77.2.0.6 | S77.GUI.2.0.5 | 2023-05-02
JR2023053 | Current Spare | Current Spare | S78B.03 | 2020-03-01 | S78B.03.0007 | S78B.3.1.1 | S78B.GUI.3.0.2
JR2023088 | Zanzibar | Globe Corporation | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0021 | AWPSIT.MC.01.1.1.8 | AWPSIT.GUI.01.1.1.9 | 2023-07-01
JR2023088 | Zanzibar | Globe Corporation | AWPSIT.MC.01 | 2018-10-01 | AWPSIT.MC.01.0022 | AWPSIT.MC.01.1.1.8 | AWPSIT.GUI.01.1.1.9 | 2023-07-01
JR2023088 | Zanzibar | Globe Corporation | ELTP.01 |  | ELTP.01.0005 | ELTP.1.0.0 | ELTP.GUI.1.0.0 | 2023-07-01 |  |  | Single Bank
JR2023088 | Zanzibar | Globe Corporation | ELTP.01 |  | ELTP.01.0006 | ELTP.1.0.0 | ELTP.GUI.1.0.0 | 2023-07-01 |  |  | Single Bank
JR2023088 | Zanzibar | Globe Corporation | VLT.02 | 2021-08-01 | VLT.02.0013 | VLT.2.0.3 | VLT.GUI.2.0.1 | 2023-07-01 |  |  | Glove and Condom.
Stock | Stock | Stock | S18C.01 |  | S18C.01.0002 | S18C.0.0.1 | S18C.GUI.0.0.1 | 2023-08-01 |  |  | First production
#VALUE! | Enersol Pty Ltd | Enersol Pty Ltd. | S18C.01 |  | S18C.01.0001 | S18C.0.0.1 | S18C.GUI.0.0.1 | 2022-05-01 |  |  | Intial prototype
#VALUE! | Inventory | - | S78.02 | 2019-04-01 | S78.02.0025 | - | - | - | -
#VALUE! |  | Jinan Huanzheng Technology & Development Co., Ltd. | S78A.03 | 2020-03-01 | S78A.03.0002 | S78A.3.1.1 | S78A.03.1.0.2 | 2023-06-22 |  |  | Replacement device as S78B was incorectly ordered initially.
#VALUE! | Spare for Malaysia | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0010 | VLT.2.0.3 | VLT.GUI.2.0.1 |  |  |  | Spare control box for malaysia to use if needed
#VALUE! | Spare for Malaysia | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0011 | VLT.2.0.3 | VLT.GUI.2.0.1 |  |  |  | Spare control box for malaysia to use if needed
#VALUE! | Spare for Malaysia | Enersol Pty Ltd. | VLT.02 | 2021-08-01 | VLT.02.0012 | VLT.2.0.3 | VLT.GUI.2.0.1 |  |  |  | Spare control box for malaysia to use if needed